## Slide 1
10 Gauge
Anterior Thoracic
Decompression
Rev A 05162023
Notes: The following program was created specifically for the ENHANCED ARS™ (Air Release System) decompression needle orientation and training.
For questions, concerns or additional information please contact:
North American Rescue®, LLC
email: info@NARescue.com
Tel: 864.675.9800
Mail: 35 Tedwall Court Greer, SC 29650-4791
Fax: 864.675.9880

## Slide 2
Our comprehensive training materials include presentation and videos for
10G & 14G Anterior & Lateral Thoracic Decompression
as well as detailed instructions for use
image from video program
image of presentation slides
Notes: Videos, Presentations, and Instructions for Use (IFUs) are part of our complete training program.

## Slide 3
Warranty: The ENHANCED ARS™ decompression needle system is a medical device, the use of which requires specific education and training. North American Rescue, LLC. warrants the ENHANCED ARS™ as merchantable expressly for the indication detailed. North American Rescue disclaims all other implied warranties relating to this product, to include use beyond this product’s identified purpose, and utilization by untrained personnel or legally unauthorized parties.
Caution: Federal Law restricts the ENHANCED ARS™ to sale by, or on the order of, a licensed physician.
the following materials were developed for the purpose of
ENHANCED ARS™ decompression needle system orientation and training
Notes: As with most modern medical devices, the ENHANCED ARS™ meets rigid safety standards, governmental mandates, and the demanding expectation of clinicians worldwide.
The ENHANCED ARS™ was developed with attention to the latest presented and published scientific evidence, as well as guidance from today’s leading experts in the fields of Emergency Medicine, Trauma Surgery, Pulmonology, Critical Care, Combat Care, and Pre-Hospital Emergency Medicine.
The ENHANCED ARS™ Warranty and Advisory should be read and understood as a component of this program.

## Slide 4
Key Term: Tension Pneumothorax. A known life-threatening medical emergency which, if left untreated, may result in death.
Indication: For relief of tension pneumothorax in the adult patient.
Contraindications: Not intended for treatment of simple pneumothorax or hemothorax. The efficacy of this device has not been established in pediatric patients.
Warning: Failure to utilize this device properly may result in injury to cardiac, pulmonary, or vascular structures.
the following materials were developed for the purpose of
ENHANCED ARS™ decompression needle system orientation and training
Notes: The Key Term, Indication and Contraindications are further reinforced within this program.
WARNING: As with all invasive medical devices, usage is predicated on proper orientation and training, careful consideration of the indications, patient condition, review of potential contraindications, and a specific RISK versus BENEFIT analysis.
Failure to utilize this device according to the directions for use, didactic materials provided, and NAR recommended hands-on training, MAY RESULT IN FURTHER INJURY.

## Slide 5
at the conclusion of didactic and hands-on training, you should be able to:
  Identify components and function
  List indications, contraindications, and expected therapeutic benefits of thoracic decompression
  Identify (right and left) lateral landmarks for thoracic decompression
  Identify (right and left) anterior landmarks for thoracic decompression
  List indications of successful thoracic decompression
  Define potential complications of improperly performed thoracic decompression procedure
  Discuss current scientific evidence as it relates to thoracic decompression
Objectives
See lateral
NDC program
Notes: At the conclusion of this specifically developed ENHANCED ARS™ didactic, and NAR recommended hands-on training program, the clinical provider should be able to:
  Identify the ENHANCED ARS™ components and function.
  List the indications, contraindications, and expected therapeutic benefits of thoracic decompression.
  Identify the (right and left) lateral landmarks for thoracic decompression. Objective is GRAY on this slide because it is covered in lateral insertion program. Suggestion – review and become competent with both insertion locations.
  Identify the (right and left) anterior landmarks for thoracic decompression.
  List the indications (signs and symptoms) of successful thoracic decompression.
  Define the potential complications of an improperly performed thoracic decompression procedure.
  Discuss current scientific evidence as it relates to thoracic decompression.
At the completion of clinical training consider offering both a written and skills competency evaluation.

## Slide 6
1
2
3
protective case (sterile if seal unbroken)
twist off cap
case cap removed (exposing proximal needle set)
Objective 1: components and function (removal from protective case)
10GA needle
10GA catheter
Notes: The ENHANCED ARS™ is packaged in a proven, robust hard-shell case, with an assurance seal indicating sterility and lot number. As with all medical devices, labeling provides additional information regarding device length and size, as well as contact information for North American Rescue®, LLC.
Instructions for use accompany EACH needle set and should be reviewed during training efforts and emergent usage when needed.
To open:
Ensure integrity of safety seal
“Twist off” red cap, breaking safety seal
Remove red case cap, exposing ENHANCED ARS™

## Slide 7: ENHANCED ARS™ decompression needle system
six (6) key system features
Objective 1: components and function
Notes: ENHANCED ARS™ decompression needle system and training materials meet and exceed 11 April 18 TCCC Guidelines Change 17-02”.
The ENHANCED ARS™ is a robustly packaged complete assembly created to address documented failures and current management challenges in the field and critical care settings.
Length: 3.25 inches (8.25cm) KEY Point: Length is necessary based on an average 2.1 inches ( 5.3 centimeters) of skin to rib depth AND an additional average of 0.5+ inches of rib to pleura depth.
Size & modifications: 10ga (10 French) catheter with distal fenestrations (holes). The fenestrations provide a 314% increase in air outflow when compared to a 14ga catheter, and a 74% increase in air outflow over a non-fenestrated 10ga.
IMPORTANT POINT: 10ga ENHANCED ARS™ catheter walls (external and internal lumen combined) have been shown to remain patent (open) during routine usage.
The catheter additionally has been designed with NVG-friendly centimeter markings along its length for precise depth control (a key factor to improve patient safety) during placement.
Needle tip has been modified to a bi-beveled cutting blade with an adjusted angle of entry to allow for a skin nick, and to improve penetration through varying types of tissue. This modification offers clinicians enhanced entry control (which prevents “explosive entry” and penetration failure, thus improving patient safety).

## Slide 8
catheter “threads” into position while needle is held stationary
Objective 1: components and function
catheter disconnects from needle by gently sliding (threading or advancing)
catheter into position
Notes: The ENHANCED ARS™
Once proper insertion depth has been achieved (not deeper than 2 cm beyond parietal pleura), and ENHANCED ARS™ has been oriented toward the mid-clavicle, thread the catheter into position while holding needle stationary.
Catheter should be threaded (advanced) into position along parietal pleural wall for maximal therapeutic benefit.

## Slide 9
Objective 1: components and function
remove entire needle (after catheter has been threaded into position)
note catheter flexibility
once removed
safely dispose of needle
in biohazard/sharps container
Notes: Note that ENHANCED ARS™ needle should be removed ONLY when the catheter has been threaded into position (an established failure point for needle decompression).
Note that 10ga catheter has improved flexibility to promote continued relief of thoracic pressure during critical patient management.

## Slide 10
Objective 1: components and function
note air escaping from thoracic cavity
fenestrations improve air release
Notes: 10ga catheter with distal fenestrations (holes). The fenestrations provide a documented increase in opportunity for air outflow (when compared to a non-fenestrated 10ga or any type of 14ga catheter).

## Slide 11
tension pneumothorax:
a life-threatening emergency, which if left untreated,
may result in death.
left tension pneumothorax
Objective 2: indications
Notes: A tension pneumothorax is known to be a life-threatening emergency, which if left untreated, may result in death.
NOTE: Illustration depicts an anterior/posterior thoracic view of patient WITH left sided tension pneumothorax.

## Slide 12
indications: consider treatment of tension pneumothorax
when one or more of the following are present:
severe or progressive respiratory distress
oxygen saturation less than 90%
shock (loss of radial pulse)
absent or markedly decreased breath sounds (if able to auscultate)
traumatic cardiac arrest without obvious fatal wounds (consider immediate placement of bilateral ENHANCED ARS™ decompression needles and consider ALL other possible causes)
Objective 2: indication
Notes: There are MANY reasons a tension pneumothorax may develop. These include both medical issues as well as trauma.
Indications: Treatment of a tension pneumothorax should be considered when one or more of the following are present:
Severe or progressive respiratory distress.
oxygen saturation less than 90%.
Shock.
Absent or markedly decreased breath sounds.
Traumatic cardiac arrest without obvious fatal wounds (consider immediate placement of bilateral ENHANCED ARS™ decompression needles and consider ALL other possible causes).

## Slide 13
Objective 2: contraindications
contraindications: the ENHANCED ARS™ is not intended for treatment of simple pneumothorax or hemothorax
efficacy of the ENHANCED ARS™ has not been established in
pediatric patients
Notes: Contraindications: The ENHANCED ARS™ is not intended for treatment of simple pneumothorax or hemothorax.
Efficacy of the ENHANCED ARS™ has not been established in pediatric patients.

## Slide 14
thoracic decompression should improve one or more of the following:
respiratory distress
relief of restrictive pressure between the parietal and visceral pleura (secondary to injury or significant medical complication)
oxygen saturation (≥ 90% may be dependent on use of supplemental oxygen)
return of radial pulse or vital signs
Objective 2: therapeutic benefits
Notes: thoracic decompression should improve one or more of the following:
Respiratory distress.
Relief of restrictive pressure between the parietal and visceral pleura (secondary to injury or significant medical complication).
Oxygen saturation (≥ 90% may be dependent on use of supplemental oxygen).
Return of radial pulse or vital signs.

## Slide 15
Objective 4: anterior insertion site
1
2
3
4
1
2
3
4
numbered
Intercostal
spaces
numbered ribs
1
2
3
4
1
2
3
4
numbered
Intercostal
spaces
numbered ribs
midclavicular line
midclavicular line
insertion site
insertion site
illustration of key anatomical structures
*insertion site
is superior to targeted rib
to avoid neuro-vascular bundle
Notes: Anterior land-marking for thoracic decompression is critical for patient safety.
Note that thoracic decompression may be relieved at the 2nd intercostal space on the midclavicular line (consult your protocol or standards for specific verification).
The illustration in this slide depicts a bilateral pneumothorax for clarity.
Looking closely you will also note that the ribs, midclavicular lines, intercostal spaces, and insertions sites (left and right) are labeled. IMPORTANT POINT: This slide depicts intercostal spaces and ribs that are numbered for clinician clarification and discussion.
Providers MUST ensure proper land-marking prior to ENHANCED ARS™ insertion.
NOTE: Insertion site is superior to the rib that is initially targeted. THIS IS A CRITICAL STEP to ensure proper placement location, avoidance of neurovascular bundle beneath each rib, and represents a known, safe starting position, from which the provider can control precise placement and insertion depth.

## Slide 16
Objective 4: anterior insertion site
Identify sternal notch
Notes: The insertion site commonly accepted for anterior decompression is the 2nd intercostal space, on the midclavicular line. Consult with local protocol or standards for verification).
Steps to locating the anterior decompression site:
1. identification of the sternal notch.

## Slide 17
Objective 4: anterior insertion site
identify sternoclavicular
joint space
Notes: The insertion site commonly accepted for anterior decompression is the 2nd intercostal space, on the midclavicular line. Please confer with protocol or standards for verification).
Steps to locating the anterior decompression site:
identification of the sternal notch.
2. Identification of the sternoclavicular joint space.

## Slide 18
Objective 4: anterior insertion site
identify acromioclavicular
ligament (notch)
Notes: The insertion site commonly accepted for anterior decompression is the 2nd intercostal space, on the midclavicular line. Please confer with protocol or standards for verification).
Steps to locating the anterior decompression site:
identification of the sternal notch.
Identification of the sternoclavicular joint space.
identification of the acromioclavicular joint space.

## Slide 19
Objective 4: anterior insertion site
simultaneously
identify sternoclavicular
joint space and
acromioclavicular
ligament (notch)
Notes: The insertion site commonly accepted for anterior decompression is the 2nd intercostal space, on the midclavicular line. Please confer with protocol or standards for verification).
Steps to locating the anterior decompression site:
identification of the sternal notch.
Identification of the sternoclavicular joint space.
identification of the acromioclavicular joint space.
Identification of the clavicles mid-point (midclavicular line)

## Slide 20
Objective 4: anterior insertion site
note middle finger
Identifying sternoclavicular
joint space and thumb on
acromioclavicular
ligament (notch)
identify middle
of clavicle
Notes: The insertion site commonly accepted for anterior decompression is the 2nd intercostal space, on the midclavicular line. Please confer with protocol or standards for verification).
Steps to locating the anterior decompression site:
identification of the sternal notch.
Identification of the sternoclavicular joint space.
identification of the acromioclavicular joint space.
Identification of the clavicles mid-point (midclavicular line)

## Slide 21
Objective 4: anterior insertion site
helpful tip: identify 2nd rib
lateral to the angle of Louis
(inferior aspect of the manubrium)
creating lateral line
from angle of Louis
places finger in
2nd intercostal space
when at the
midclavicular line
Notes: The insertion site commonly accepted for anterior decompression is the 2nd intercostal space, on the midclavicular line. Please confer with protocol or standards for verification).
Steps to locating the anterior decompression site:
identification of the sternal notch.
Identification of the sternoclavicular joint space.
identification of the acromioclavicular joint space.
Identification of the clavicles mid-point (and midclavicular line)
Identification of the 2nd intercostal space (above the third rib - midclavicular line)
    Helpful tip: Identify 2nd rib lateral to the angle of Louis

## Slide 22
Objective 4: anterior insertion site
helpful tip: creating lateral line
from angle of Louis
places finger on
2nd intercostal space
site verification is critical
do not insert decompression
needle medial to midclavicular line
Notes: The insertion site commonly accepted for anterior decompression is the 2nd intercostal space, on the midclavicular line. Please confer with protocol or standards for verification).
Steps to locating the anterior decompression site:
identification of the sternal notch.
Identification of the sternoclavicular joint space.
identification of the acromioclavicular joint space.
Identification of the clavicles mid-point (and midclavicular line)
Identification of the 2nd intercostal space (above the third rib – midclavicular line)
    Helpful tip: Identify 2nd rib lateral to the angle of Louis
Identify the intersection between vertical midclavicular line and a horizontal line running laterally from the angle of Louis.

## Slide 23
Objective 4: anterior insertion site
helpful tip: compress firmly
to ensure identification of 3rd rib
mid-clavicular line
mark and cleanse
area over 3rd rib
mid-clavicular line
Notes: The insertion site commonly accepted for anterior decompression is the 2nd intercostal space, on the midclavicular line. Please confer with protocol or standards for verification).
Steps to locating the anterior decompression site:
identification of the sternal notch.
Identification of the sternoclavicular joint space.
identification of the acromioclavicular joint space.
Identification of the clavicles mid-point (and midclavicular line)
Identification of the 2nd intercostal space (above the third rib – midclavicular line)
    Helpful tip: Identify 2nd rib lateral to the angel of Louis
Identify the intersection between the vertical midclavicular line and a horizontal line running laterally from the angle of Louis.
Confirm location by firmly palpating the 3rd rib on the midclavicular line.
Cleanse site according to local protocol and standard.

## Slide 24
Objective 4: anterior insertion site
decompression needle set in position
note hand positioning to prevent needle and catheter separation
3rd rib on mid-clavicular
line re-identified
site verification
with framing
consider nicking skin with needle tip
to ease ENHANCED ARS™ penetration
through epidermis and dermis
Notes: Once site has been properly cleansed, preparations for ENHANCED ARS™ insertion can be completed.
Note site “framing” and “reverification” in the image provided. Framing (spreading the skin between two fingers) tightens the tissue to improve penetration.
Site reverification (of the targeted rib) ensures that the ENHANCED ARS™ needle tip will FIRST be placed against the rib.
Consider nicking skin with needle tip to ease ENHANCED ARS™ penetration through epidermis and dermis.

## Slide 25
Objective 4: anterior insertion site
needle set lined up
to penetrate skin
and identify 3rd rib
ENHANCED ARS™ in position
and perpendicular to
chest wall
Notes: Once the ENHANCED ARS™ has penetrated the skin, adipose, and muscle tissue, it should be firmly placed against the rib (this is a critical identification step to ensure patient safety).
Once the ENHANCED ARS™ needle tip has been placed against the targeted rib, the provider should MARK and HOLD 3 centimeters above the skin.
Depth marking is critical to the next phase of insertion (which is placement of the needle set into the pleural space).
ENSURE that the SPEAR™ is positioned perpendicular to the chest wall. Note the sloped angle of the anterior thoracic wall.

## Slide 26
Objective 4: anterior insertion site
mid-clavicular line
Perpendicular to chest wall
needle set stabilized between
fingers, hub resting on palm
needle set penetrating skin and identifying 3rd rib
needle set stabilized between
thumb and middle finger
hub held by index finger
Notes: Once the ENHANCED ARS™ has penetrated the skin and adipose tissue, it should be firmly placed against the rib (this is a critical identification step to ensure patient safety).
Once the ENHANCED ARS™ needle tip has been placed against the targeted rib, the provider should MARK and HOLD 3 centimeters above the skin.
Depth marking is critical to the next phase of insertion (which is placement of the needle set into the pleural space).
ENSURE that the SPEAR™ is positioned perpendicular to the chest wall. Note the sloped angle of the anterior thoracic wall.

## Slide 27
Objective 4: anterior insertion
needle set stabilization
with fingertip depth control
insertion site is
superior to targeted rib
needle set lined up
to penetrate skin
and identify 3rd rib
Notes: Illustration on this slide is for additional clarity of previous point.
Once the ENHANCED ARS™ has penetrated the skin and adipose tissue, it should be firmly placed against the rib (this is a critical identification step to ensure patient safety).
Once the ENHANCED ARS™ needle tip has been placed against the targeted rib, the provider should MARK and HOLD 3 centimeters above the skin.
Depth marking is critical to the next phase of insertion (which is placement of the needle set into the pleural space).

## Slide 28
Objective 4: anterior insertion site
note needle stabilization
and fingertip depth control
Insert NO MORE THAN
2 cm into pleural cavity
DO NOT “BURY” OR “HUB” NEEDLE SET
Notes: Once the ENHANCED ARS™ needle tip has been placed against the targeted rib, the provider should MARK and HOLD 3 centimeters above the skin.
Proceed with insertion while ensuring that penetration is limited by fingertip control.
Depth marking is critical to the next phase of insertion (which is placement of the needle set into the pleural space).
ENSURE that the ENHANCED ARS™ is positioned perpendicular to the chest wall. Note the sloped angle of the anterior thoracic wall.

## Slide 29
insert ENHANCED ARS™ superiorly
over targeted rib
limit insertion 2 cm into pleural cavity
KEY POINT:
stabilization and control
target rib and precisely
penetrate chest wall to
prevent further injury
Objective 4: anterior insertion site
Notes: Illustration on this slide is for additional clarity of previous point.
With 3 centimeters identified, the provider must carefully move the ENHANCED ARS™ needle tip superiorly over the targeted rib and guide the assembly into the thoracic cavity. Illustrations on this slide depict needle set entry over the targeted rib. Note that this image shows the ENHANCED ARS™ advanced the correct distance with “fingertip control.”
This step was safely accomplished by pinning 3 centimeters above the tissue WHILE the needle tip was against the rib. Note that the provider ASSURED proper insertion depth was achieved (and NOT exceeded).
Critical Point: Ensure that ENHANCED ARS™ placement is accomplished with the needle set perpendicular to the thoracic wall.
It is well understood and established that some experienced providers can appreciate needle set penetration of the parietal pleura (thus indicating appropriately achieved depth). Regardless of depth control method - DO NOT “BURY” or fully extend ANY decompression needle completely into the thoracic cavity.
*Air under pressure may release from the thoracic cavity at this point.

## Slide 30
needle MUST remain stationary
while threading catheter
Objective 4: anterior insertion site
thread catheter
toward middle of clavicle
direct catheter
toward middle of clavicle
Notes: Once catheter has been released from the needle, advance the catheter while keeping the needle stationary as a guide.
*Air under pressure may release from the thoracic cavity at this point.

## Slide 31
Objective 4: anterior insertion site
completely thread catheter into pleural cavity
needle should remain stationary
needle should remain stationary
Notes: Illustration on this slide is for additional clarity of previous point.
Once catheter has been released from the needle, advance the catheter while keeping the needle stationary as a catheter guide.
*Air under pressure may release from the thoracic cavity at this point.

## Slide 32
Objective 4: anterior insertion site
secure catheter according to
local protocol and standards
. . . . .
relief of tension
Notes: Secure the catheter according to local protocol and standard.
Monitor the patient.

## Slide 33
successful thoracic decompression may have occurred if one or more of the following is observed:
improvement of respiratory distress
relief of air from catheter
improvement of oxygen saturation (≥ 90% may be dependent on use of supplemental oxygen)
return of radial pulse or vital signs
Objective 5: indications of successful thoracic decompression
Notes: Successful thoracic decompression may have occurred if one or more of the following is observed:
Improvement of respiratory distress.
Relief of air from catheter.
Improvement of oxygen saturation (≥ 90% may be dependent on use of supplemental oxygen).
Return of radial pulse or vital signs.

## Slide 34
following thoracic decompression procedure, continually assess patient for complications:
hemodynamic instability
respiratory distress
unilateral chest expansion
decreased oxygen saturation
bleeding
catheter occlusion
hematoma
if two needle decompression attempts fail to relieve condition consider other causes and potential treatments
IS THIS A CIRCULATORY PROBLEM?
Objective 6: complications
Notes: A patient should be monitored following any medical procedure. Thoracic decompression (by definition) is a life-saving procedure that is NOT without risk.
With continuous training, as well as safeguards and careful considerations, patient and provider risk can be mitigated.
Following thoracic decompression, continually assess the patient for complications. These should include (but not be limited to):
Hemodynamic instability
Respiratory distress
Unilateral chest expansion
Decreased oxygen saturation
Bleeding
Catheter occlusion
Hematoma
If two needle decompression attempts fail to relieve the suspected problem, consider other causes and potential treatments. IS THIS A CIRCULATORY PROBLEM?

## Slide 35
potential adverse complications: of improper thoracic decompression
death secondary to cardiac penetration
lung injury
vascular injury
nerve damage
  pain
  numbness
  paralysis of intercostal muscle
infection
Objective 6: complications
Notes: As with all invasive medical devices, potential adverse complications are a risk that MUST be considered. Most often, complications are the result of improperly performed interventions. Thoracic decompression is NOT WITHOUT RISK.
The FIRST step to avoiding injury is a complete understanding of the indications and contraindications for ENHANCED ARS™ usage. Each of the listed complications SHOULD be discussed WITH specific mitigation strategies integrated into the training program.
Complications are bolded in BLACK
Mitigation strategies presented in GRAY
Potential complications include (but are not limited to):
Death secondary to cardiac penetration (clinician should control thoracic penetration and restrict depth of needle set insertion)
Lung injury (insertion of needle set should be 2 cm past the parietal pleura)
Vascular injury (1. insertion site restrictions to the specific locations identified in this training program. 2. maintain proper needle set angle during insertion.
3. avoidance of needle set over penetration.
Nerve damage (ensure proper insertion site location “superior to targeted rib”)
  Pain (if time permits, consider local anesthetic per protocol or standard)
  Numbness (avoid insertion on the inferior aspect of the rib)
  Paralysis of intercostal muscle (avoid insertion below rib)
Infection (use aseptic technique)

## Slide 36
Objective 7: scientific evidence
The ENHANCED ARS™ decompression needle system, and this presentation, were developed utilizing the latest published evidence, independent research, and the support of dedicated Military and Civilian medical professionals in Emergency Medicine, Trauma Surgery, Pulmonology, Radiology, and Pathology.
Clinical providers, regardless of their position, must dedicate themselves to the unrelenting truth that critical care is an evolution on behalf of those in need.
Butler F, Holcomb J, Shackelford S, et al. Management of the Suspected Tension Pneumothorax in Tactical Combat Care, TCCC Guidelines Change 17-02. J Spe Op Med. 2018; 18: 19-35.
*The aforementioned publication references ninety-six additional papers worthy of careful review.
Kruse A, Achay J, et al Cross-Over Study Comparing 14G and 10G Needle Decompression in a Cadaver Pneumothorax Model SOMA Training, Education & Scientific Assembly 2023, Presentation.
Notes: The ENHANCED ARS™ decompression needle system, and this presentation, were developed utilizing the latest published evidence, independent research, and the support of dedicated Military and Civilian medical professionals in Emergency Medicine, Trauma Surgery, Pulmonology, Radiology, and Pathology.
Clinical providers, regardless of their position, must dedicate themselves to the unrelenting truth that critical care is an evolution on behalf of those in need.
Butler F, Holcomb J, Shackelford S, et al. Management of the Suspected Tension Pneumothorax in Tactical Combat Care, TCCC Guidelines Change 17-02. J Spe Op Med. 2018; 18: 19-35.
*The aforementioned publication references ninety-six additional papers worthy of careful review.
Kruse A, Achay J, et al Cross-Over Study Comparing 14G and 10G Needle Decompression in a Cadaver Pneumothorax Model SOMA Training, Education & Scientific Assembly 2023, Presentation.

## Slide 37
Butler F, Holcomb J, Shackelford S, et al. Management of the Suspected Tension Pneumothorax in Tactical Combat Care, TCCC Guidelines Change 17-02. J Spe Op Med. 2018; 18: 19-35.
Nine Key Facets of Tactical Combat Care Guidelines Change 17-02
Continuation of aggressive approach to suspecting and treating tension pneumothorax
Emphasis of bilateral decompression in traumatic arrest
Addition of 10 Gauge catheter
Designates either Lateral or Anterior sites as acceptable for thoracic decompression
Addition of procedural elements (critical procedural modifications are included within this material)
Defines successful thoracic decompression
Recommends ONLY two needle decompressions be attempted before moving on to circulation
Addition of materials that recommend consideration of tension pneumothorax in presentations of shock
Addition of simple thoracostomy (if presentation warrants - following two unsuccessful needle decompression attempts - and provider is trained)
Objective 7: scientific evidence
Notes: There are nine Key Facets to thoracic decompression in the Tactical Combat Care Guidelines Change 17-02 that should be considered: Slide numbers correspond to Anterior Insertion Program
Continuation of aggressive approach to suspecting and treating tension pneumothorax.
There is discussion and some evidence that unwarranted needle decompressions have presented complications. While this is a valid concern, failure to recognize and treat tension pneumothorax remains a lethal reality.
Emphasis of bilateral decompression in traumatic arrest.
Failure to completely address ALL possible causes of traumatic arrest continues to present as a lethal concern.
Adds 10 Gauge.
Designates either Lateral or Anterior sites as acceptable for thoracic decompression.
Defines successful thoracic decompression.
Recommends ONLY two needle decompressions be attempted before moving on to circulation.
Addition of materials that recommend consideration of tension pneumothorax in presentations of shock.
Addition of finger thoracostomy when presentation warrants. Provider has attempted two needle decompressions without success and is authorized to perform procedure.

## Slide 38
For additional information
about the ENHANCED ARS™
email: info@NARescue.com
Tel: 864.675.9800
Mail: 35 Tedwall Court Greer, SC 29650-4791
Fax: 864.675.9880
Notes: Thank you for participating in the ENHANCED ARS™ (Air Release System) decompression needle orientation and training.
If you have any questions, concerns or would like additional information please contact:
North American Rescue®, LLC
email: info@NARescue.com
Tel: 864.675.9800
Mail: 35 Tedwall Court Greer, SC 29650-4791
Fax: 864.675.9880